## Perf test 20230905
Manufacturer | James G. Biddle Co
Model | 601147-1
Description | Calibration Resistance Standard
Serial No. | 43056
Accuracy specification | ± | 0.03 | % of indication
| ± | 0.0005 | Ω for each decade involved
Notes | All measurements were made using hp34401a 6.5 digit DMM after 1hr warmup. Measurements are intended to document the performance of the instrument at the time of testing. Results should not be considered a valid metrological calibration. Results are not guaranteed to remain the same at any time after the testing was performed.
Test date:09/05/2023
|  |  | Nominal (Ω) | Decade positions | [7 empty] | Performance Specifications |  | Measured |  |  |  | Uncertainty |  | Measured (Ω) (after zero subtraction) | %Error | Abs. Error (Ω) | Pass/Fail | P/F low est | P/F high est
|  |  |  | x100k | x10k | x1k | x100 | x10 | x1 | x.1 | x.01 | Lower Limit (Ω) | Upper Limit (Ω) | Mean (Ω) | Min (Ω) | Max (Ω) | Count | Lower Limit | Upper Limit
|  |  | 000000.00 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0.0089222 | 0.00802 | 0.009551 | 100 |  |  | 0 | NA | NA | Pass
|  |  | 0.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0.009497 | 0.010503 | 0.019271 | 0.018248 | 0.020633 | 100 | 0.0091565 | 0.011541 | 0.010349 | 3.4895 | 0.00034895 | Pass | Fail | Fail
|  |  | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0.019494 | 0.020506 | 0.029354 | 0.028489 | 0.029706 | 100 | 0.019823 | 0.02104 | 0.020432 | 2.1593 | 0.00043186 | Pass | Pass | Fail
|  |  | 0.03 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0.029491 | 0.030509 | 0.039354 | 0.038241 | 0.040437 | 100 | 0.029333 | 0.031529 | 0.030431 | 1.4382 | 0.00043146 | Pass | Fail | Fail
|  |  | 0.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0.039488 | 0.040512 | 0.049318 | 0.04882 | 0.049686 | 100 | 0.039962 | 0.040828 | 0.040395 | 0.9883 | 0.00039532 | Pass | Pass | Fail
|  |  | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0.049485 | 0.050515 | 0.059257 | 0.058836 | 0.05999 | 100 | 0.049758 | 0.050912 | 0.050335 | 0.6701 | 0.00033505 | Pass | Pass | Fail
|  |  | 0.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 0.059482 | 0.060518 | 0.069125 | 0.0686 | 0.069516 | 100 | 0.059745 | 0.060661 | 0.060203 | 0.3387 | 0.00020322 | Pass | Pass | Fail
|  |  | 0.07 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 0.069479 | 0.070521 | 0.079048 | 0.078088 | 0.079518 | 100 | 0.069411 | 0.070841 | 0.070126 | 0.1794 | 0.00012558 | Pass | Fail | Fail
|  |  | 0.08 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0.079476 | 0.080524 | 0.088938 | 0.088291 | 0.089659 | 100 | 0.079332 | 0.0807 | 0.080016 | 0.020137 | 1.611e-05 | Pass | Fail | Fail
|  |  | 0.09 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9 | 0.089473 | 0.090527 | 0.098828 | 0.098243 | 0.099323 | 100 | 0.089365 | 0.090445 | 0.089905 | -0.10502 | -9.452e-05 | Pass | Fail | Pass
|  |  | 0.1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 0.09947 | 0.10053 | 0.10882 | 0.10772 | 0.10934 | 100 | 0.099087 | 0.10071 | 0.099896 | -0.10379 | -0.00010379 | Pass | Fail | Fail
|  |  | 0.1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0.09947 | 0.10053 | 0.10882 | 0.10826 | 0.10954 | 100 | 0.099255 | 0.10054 | 0.099895 | -0.10479 | -0.00010479 | Pass | Fail | Fail
|  |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0.19944 | 0.20056 | 0.20877 | 0.20822 | 0.21035 | 100 | 0.19878 | 0.20092 | 0.19985 | -0.075245 | -0.00015049 | Pass | Fail | Fail
|  |  | 0.3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0.29941 | 0.30059 | 0.30879 | 0.30776 | 0.30914 | 100 | 0.29918 | 0.30056 | 0.29987 | -0.04373 | -0.00013119 | Pass | Fail | Pass
|  |  | 0.4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0.39938 | 0.40062 | 0.40878 | 0.40817 | 0.40931 | 100 | 0.39929 | 0.40043 | 0.39986 | -0.035098 | -0.00014039 | Pass | Fail | Pass
|  |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0.49935 | 0.50065 | 0.50894 | 0.50803 | 0.51008 | 100 | 0.499 | 0.50104 | 0.50002 | 0.004042 | 2.021e-05 | Pass | Fail | Fail
|  |  | 0.6 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0.59932 | 0.60068 | 0.60864 | 0.60748 | 0.60934 | 100 | 0.59878 | 0.60064 | 0.59971 | -0.047798 | -0.00028679 | Pass | Fail | Pass
|  |  | 0.7 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | 0.69929 | 0.70071 | 0.70907 | 0.70798 | 0.70949 | 100 | 0.69939 | 0.7009 | 0.70014 | 0.02043 | 0.00014301 | Pass | Pass | Fail
|  |  | 0.8 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0.79926 | 0.80074 | 0.80867 | 0.80807 | 0.809 | 100 | 0.79928 | 0.80021 | 0.79975 | -0.031536 | -0.00025229 | Pass | Pass | Pass
|  |  | 0.9 | 0 | 0 | 0 | 0 | 0 | 0 | 9 | 0 | 0.89923 | 0.90077 | 0.90884 | 0.90856 | 0.90927 | 100 | 0.89957 | 0.90027 | 0.89992 | -0.0092767 | -8.349e-05 | Pass | Pass | Pass
|  |  | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 0 | 0.9992 | 1.0008 | 1.0088 | 1.0085 | 1.0091 | 100 | 0.99955 | 1.0002 | 0.99989 | -0.010919 | -0.00010919 | Pass | Pass | Pass
|  |  | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0.9992 | 1.0008 | 1.0089 | 1.0082 | 1.0095 | 100 | 0.99935 | 1.0006 | 0.99999 | -0.000819 | -8.19e-06 | Pass | Pass | Pass
|  |  | 2 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1.9989 | 2.0011 | 2.009 | 2.0084 | 2.0094 | 100 | 1.9995 | 2.0006 | 2 | 0.0019405 | 3.881e-05 | Pass | Pass | Pass
|  |  | 3 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 2.9986 | 3.0014 | 3.0091 | 3.0087 | 3.0095 | 100 | 2.9998 | 3.0006 | 3.0002 | 0.0058937 | 0.00017681 | Pass | Pass | Pass
|  |  | 4 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 3.9983 | 4.0017 | 4.0091 | 4.0087 | 4.0094 | 100 | 3.9998 | 4.0005 | 4.0001 | 0.0033453 | 0.00013381 | Pass | Pass | Pass
|  |  | 5 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 4.998 | 5.002 | 5.0092 | 5.0078 | 5.0099 | 100 | 4.9992 | 5.0013 | 5.0002 | 0.0046562 | 0.00023281 | Pass | Pass | Pass
|  |  | 6 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 5.9977 | 6.0023 | 6.0091 | 6.0086 | 6.0095 | 100 | 5.9997 | 6.0006 | 6.0002 | 0.0028302 | 0.00016981 | Pass | Pass | Pass
|  |  | 7 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | 0 | 6.9974 | 7.0026 | 7.009 | 7.0081 | 7.0095 | 100 | 6.9994 | 7.0007 | 7 | 0.000483 | 3.381e-05 | Pass | Pass | Pass
|  |  | 8 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 7.9971 | 8.0029 | 8.0089 | 8.0083 | 8.0092 | 100 | 7.9995 | 8.0004 | 8 | -0.00025237 | -2.019e-05 | Pass | Pass | Pass
|  |  | 9 | 0 | 0 | 0 | 0 | 0 | 9 | 0 | 0 | 8.9968 | 9.0032 | 9.0088 | 9.0078 | 9.0101 | 100 | 8.9987 | 9.001 | 8.9999 | -0.0015021 | -0.00013519 | Pass | Pass | Pass
|  |  | 10 | 0 | 0 | 0 | 0 | 0 | 10 | 0 | 0 | 9.9965 | 10.0035 | 10.0089 | 10.008 | 10.0093 | 100 | 9.9993 | 10.0006 | 9.9999 | -0.0007219 | -7.219e-05 | Pass | Pass | Pass
|  |  | 10 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 9.9965 | 10.0035 | 10.0083 | 10.0077 | 10.0087 | 100 | 9.9988 | 9.9999 | 9.9993 | -0.0065219 | -0.00065219 | Pass | Pass | Pass
|  |  | 20 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 19.9935 | 20.0065 | 20.0075 | 20.0069 | 20.008 | 100 | 19.998 | 19.9991 | 19.9985 | -0.007311 | -0.0014622 | Pass | Pass | Pass
|  |  | 30 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 29.9905 | 30.0095 | 30.0066 | 30.0059 | 30.007 | 100 | 29.9971 | 29.9982 | 29.9976 | -0.0079073 | -0.0023722 | Pass | Pass | Pass
|  |  | 40 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 39.9875 | 40.0125 | 40.0059 | 40.0049 | 40.0063 | 100 | 39.9963 | 39.9976 | 39.997 | -0.0075805 | -0.0030322 | Pass | Pass | Pass
|  |  | 50 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 49.9845 | 50.0155 | 50.0051 | 50.0044 | 50.0057 | 100 | 49.9955 | 49.9969 | 49.9962 | -0.0076444 | -0.0038222 | Pass | Pass | Pass
|  |  | 60 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 59.9815 | 60.0185 | 60.004 | 60.0024 | 60.0046 | 100 | 59.994 | 59.9962 | 59.9951 | -0.0081703 | -0.0049022 | Pass | Pass | Pass
|  |  | 70 | 0 | 0 | 0 | 0 | 7 | 0 | 0 | 0 | 69.9785 | 70.0215 | 70.0033 | 70.0023 | 70.0042 | 100 | 69.9934 | 69.9953 | 69.9944 | -0.008046 | -0.0056322 | Pass | Pass | Pass
|  |  | 80 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 79.9755 | 80.0245 | 80.0022 | 80.0012 | 80.0027 | 100 | 79.9926 | 79.9941 | 79.9933 | -0.0083652 | -0.0066922 | Pass | Pass | Pass
|  |  | 90 | 0 | 0 | 0 | 0 | 9 | 0 | 0 | 0 | 89.9725 | 90.0275 | 90.0014 | 90.0005 | 90.0028 | 100 | 89.9914 | 89.9937 | 89.9925 | -0.0083247 | -0.0074922 | Pass | Pass | Pass
|  |  | 100 | 0 | 0 | 0 | 0 | 10 | 0 | 0 | 0 | 99.9695 | 100.0305 | 100.0008 | 100.0001 | 100.0014 | 100 | 99.9912 | 99.9925 | 99.9919 | -0.0081222 | -0.0081222 | Pass | Pass | Pass
|  |  | 100 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 99.9695 | 100.0305 | 100.0061 | 100.0052 | 100.0069 | 100 | 99.9963 | 99.998 | 99.9972 | -0.0028222 | -0.0028222 | Pass | Pass | Pass
|  |  | 200 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 199.9395 | 200.0605 | 200.0154 | 200.0148 | 200.0166 | 100 | 200.0056 | 200.0074 | 200.0065 | 0.0032389 | 0.0064778 | Pass | Pass | Pass
|  |  | 300 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 299.9095 | 300.0905 | 300.0147 | 300.0139 | 300.0153 | 100 | 300.0051 | 300.0065 | 300.0058 | 0.0019259 | 0.0057778 | Pass | Pass | Pass
|  |  | 400 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 399.8795 | 400.1205 | 400.015 | 400.0142 | 400.016 | 100 | 400.0052 | 400.007 | 400.0061 | 0.0015195 | 0.0060778 | Pass | Pass | Pass
|  |  | 500 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 499.8495 | 500.1505 | 500.0156 | 500.0144 | 500.0168 | 100 | 500.0055 | 500.0079 | 500.0067 | 0.0013356 | 0.0066778 | Pass | Pass | Pass
|  |  | 600 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 599.8195 | 600.1805 | 600.0143 | 600.0133 | 600.015 | 100 | 600.0045 | 600.0062 | 600.0054 | 0.0008963 | 0.0053778 | Pass | Pass | Pass
|  |  | 700 | 0 | 0 | 0 | 7 | 0 | 0 | 0 | 0 | 699.7895 | 700.2105 | 700.0142 | 700.0134 | 700.0157 | 100 | 700.0041 | 700.0064 | 700.0053 | 0.00075397 | 0.0052778 | Pass | Pass | Pass
|  |  | 800 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 0 | 799.7595 | 800.2405 | 800.0115 | 800.0109 | 800.0124 | 100 | 800.0018 | 800.0033 | 800.0026 | 0.00032223 | 0.0025778 | Pass | Pass | Pass
|  |  | 900 | 0 | 0 | 0 | 9 | 0 | 0 | 0 | 0 | 899.7295 | 900.2705 | 900.014 | 900.0132 | 900.0153 | 100 | 900.004 | 900.0061 | 900.0051 | 0.0005642 | 0.0050778 | Pass | Pass | Pass
|  |  | 1000 | 0 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 999.6995 | 1000.3005 | 1000.014 | 1000.013 | 1000.014 | 100 | 1000.0046 | 1000.0056 | 1000.0051 | 0.00050778 | 0.0050778 | Pass | Pass | Pass
|  |  | 1000 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 999.6995 | 1000.3005 | 1000.022 | 1000.021 | 1000.023 | 100 | 1000.0121 | 1000.0141 | 1000.0131 | 0.0013078 | 0.013078 | Pass | Pass | Pass
|  |  | 2000 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1999.3995 | 2000.6005 | 2000.152 | 2000.144 | 2000.16 | 100 | 2000.1351 | 2000.1511 | 2000.1431 | 0.0071539 | 0.14308 | Pass | Pass | Pass
|  |  | 3000 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 2999.0995 | 3000.9005 | 3000.192 | 3000.182 | 3000.205 | 100 | 3000.1716 | 3000.1946 | 3000.1831 | 0.0061026 | 0.18308 | Pass | Pass | Pass
|  |  | 4000 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 3998.7995 | 4001.2005 | 4000.19 | 4000.181 | 4000.226 | 100 | 4000.1586 | 4000.2036 | 4000.1811 | 0.0045269 | 0.18108 | Pass | Pass | Pass
|  |  | 5000 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 4998.4995 | 5001.5005 | 5000.156 | 5000.146 | 5000.166 | 100 | 5000.1371 | 5000.1571 | 5000.1471 | 0.0029416 | 0.14708 | Pass | Pass | Pass
|  |  | 6000 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 5998.1995 | 6001.8005 | 6000.136 | 6000.126 | 6000.148 | 100 | 6000.1161 | 6000.1381 | 6000.1271 | 0.002118 | 0.12708 | Pass | Pass | Pass
|  |  | 7000 | 0 | 0 | 7 | 0 | 0 | 0 | 0 | 0 | 6997.8995 | 7002.1005 | 7000.116 | 7000.073 | 7000.158 | 100 | 7000.0646 | 7000.1496 | 7000.1071 | 0.0015297 | 0.10708 | Pass | Pass | Pass
|  |  | 8000 | 0 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 7997.5995 | 8002.4005 | 8000.09 | 8000.064 | 8000.101 | 100 | 8000.0626 | 8000.0996 | 8000.0811 | 0.0010135 | 0.081078 | Pass | Pass | Pass
|  |  | 9000 | 0 | 0 | 9 | 0 | 0 | 0 | 0 | 0 | 8997.2995 | 9002.7005 | 9000.063 | 9000.036 | 9000.081 | 100 | 9000.0316 | 9000.0766 | 9000.0541 | 0.00060086 | 0.054078 | Pass | Pass | Pass
|  |  | 10000 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 9996.9995 | 10003.0005 | 10000.04 | 10000.02 | 10000.07 | 100 | 10000.0061 | 10000.0561 | 10000.0311 | 0.00031078 | 0.031078 | Pass | Pass | Pass
|  |  | 10000 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 9996.9995 | 10003.0005 | 9998.644 | 9998.634 | 9998.671 | 100 | 9998.6166 | 9998.6536 | 9998.6351 | -0.013649 | -1.3649 | Pass | Pass | Pass
|  |  | 20000 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 19993.9995 | 20006.0005 | 19997.55 | 19997.17 | 19997.77 | 100 | 19997.2411 | 19997.8411 | 19997.5411 | -0.012295 | -2.4589 | Pass | Pass | Pass
|  |  | 30000 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 29990.9995 | 30009.0005 | 29996.77 | 29996.28 | 29997.11 | 100 | 29996.3461 | 29997.1761 | 29996.7611 | -0.010796 | -3.2389 | Pass | Pass | Pass
|  |  | 40000 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 39987.9995 | 40012.0005 | 39995.52 | 39995.11 | 39997.14 | 100 | 39994.4961 | 39996.5261 | 39995.5111 | -0.011222 | -4.4889 | Pass | Pass | Pass
|  |  | 50000 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 49984.9995 | 50015.0005 | 49995.42 | 49994.31 | 49996.5 | 100 | 49994.3161 | 49996.5061 | 49995.4111 | -0.0091778 | -4.5889 | Pass | Pass | Pass
|  |  | 60000 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 59981.9995 | 60018.0005 | 59994.53 | 59994.01 | 59996.16 | 100 | 59993.4461 | 59995.5961 | 59994.5211 | -0.0091315 | -5.4789 | Pass | Pass | Pass
|  |  | 70000 | 0 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 69978.9995 | 70021.0005 | 69993.57 | 69992.02 | 69994.74 | 100 | 69992.2011 | 69994.9211 | 69993.5611 | -0.0091985 | -6.4389 | Pass | Pass | Pass
|  |  | 80000 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 79975.9995 | 80024.0005 | 79992.69 | 79990.42 | 79994.74 | 100 | 79990.5211 | 79994.8411 | 79992.6811 | -0.0091487 | -7.3189 | Pass | Pass | Pass
|  |  | 90000 | 0 | 9 | 0 | 0 | 0 | 0 | 0 | 0 | 89972.9995 | 90027.0005 | 89991.87 | 89990.41 | 89995.6 | 100 | 89989.2661 | 89994.4561 | 89991.8611 | -0.0090432 | -8.1389 | Pass | Pass | Pass
|  |  | 100000 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 99969.9995 | 100030.0005 | 99990.77 | 99990.18 | 99992.81 | 100 | 99989.4461 | 99992.0761 | 99990.7611 | -0.0092389 | -9.2389 | Pass | Pass | Pass
|  |  | 100000 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 99969.9995 | 100030.0005 | 100005.6 | 100004.7 | 100009 | 100 | 100003.4411 | 100007.7411 | 100005.5911 | 0.0055911 | 5.5911 | Pass | Pass | Pass
|  |  | 200000 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 199939.9995 | 200060.0005 | 200046.3 | 200038.4 | 200058.4 | 100 | 200036.2911 | 200056.2911 | 200046.2911 | 0.023146 | 46.2911 | Pass | Pass | Pass
|  |  | 300000 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 299909.9995 | 300090.0005 | 300066.3 | 300061.1 | 300070.6 | 100 | 300061.5411 | 300071.0411 | 300066.2911 | 0.022097 | 66.2911 | Pass | Pass | Pass
|  |  | 400000 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 399879.9995 | 400120.0005 | 400087 | 400079.6 | 400108.4 | 100 | 400072.5911 | 400101.3911 | 400086.9911 | 0.021748 | 86.9911 | Pass | Pass | Pass
|  |  | 500000 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 499849.9995 | 500150.0005 | 500084.4 | 500059.9 | 500092.9 | 100 | 500067.8911 | 500100.8911 | 500084.3911 | 0.016878 | 84.3911 | Pass | Pass | Pass
|  |  | 600000 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 599819.9995 | 600180.0005 | 600076.3 | 600062.8 | 600107.8 | 100 | 600053.7911 | 600098.7911 | 600076.2911 | 0.012715 | 76.2911 | Pass | Pass | Pass
|  |  | 700000 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 699789.9995 | 700210.0005 | 700077.8 | 700057.2 | 700122.9 | 100 | 700044.9411 | 700110.6411 | 700077.7911 | 0.011113 | 77.7911 | Pass | Pass | Pass
|  |  | 800000 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 799759.9995 | 800240.0005 | 800073.7 | 800058.6 | 800109.7 | 100 | 800048.1411 | 800099.2411 | 800073.6911 | 0.0092114 | 73.6911 | Pass | Pass | Pass
|  |  | 900000 | 9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 899729.9995 | 900270.0005 | 900070.5 | 900044.7 | 900135.7 | 100 | 900024.9911 | 900115.9911 | 900070.4911 | 0.0078323 | 70.4911 | Pass | Pass | Pass
|  |  | 1000000 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 999699.9995 | 1000300.0005 | 1000061 | 1000045 | 1000084 | 100 | 1000041.4911 | 1000080.4911 | 1000060.9911 | 0.0060991 | 60.9911 | Pass | Pass | Pass